# Signature 55Hz – Postprocessing guidelines

Link to this guide:

Signature 55Hz processing guide.docx

### Last update: 17/08/2026 10:28:07

- RAW-file: use for post-processing

- Averaged file: Quick check for onboard (plot STAGE 1 check list)

- Current git branch:

https://github.com/ScotMarPhys/m_moorproc_toolbox/main/

- Other matlab toolboxes needed: rodb, GSW v3.06

## QC flags:

- 0: QC_NOT_EVALUATED

- 1: QC_GOOD

- 2: QC_UNKNOWN

- 3: QC_PROBABLY_BAD

- 4: QC_BAD

- 5: QC_CHANGED

- 6: QC_UNSAMPLED

- 7: QC_INTERPOLATED

- 8: QC_?

- 9: QC_MISSING

## STAGE 0 – Download data

Download raw data (*.ad2cp and *avgd.ad2cp) from instrument using the Recorder Data Retrieval tab in Nortek Signature Deployment software. Open ad2cp-files in Signature Viewer and export to mat-files (*.mat, *avgd.mat). This will automatically apply a quality control with only data where the correlation of echo and signal is above 50% is kept in the avgd.mat file.

Please ensure to rename the files in line with OSNAP processing toolbox:

Naming conventions: [serial_number]_data.mat, [seriel_number]_data_avgd.mat

E.g. RHADCP has serial number 20044: 20044_data.mat; 20044_data_avgd.mat

Copy to folder: \osnap\moorings\raw\cruise_id\adcp_s55\

E.g.: osnap\moorings\raw\jc238\adcp_s55\

If not yet done for the microcat, create a data processing control file [moorname]info.dat for the mooring in the directory osnap/moorings/proc/[moorname]. [moorname]info.dat contains metadata for mooring position (lon, lat, waterdepth and mean magnetic deviation during the deployment), deployment period, nominal depths and serial numbers of each instrument.

Example for JC238, mooring deployment #1 for RHADCP from 2020-2022:

moorname = rhadcp_01_2020

Example: osnap\moorings\proc\rhadcp_01_2020\rhadcp_01_2020info.dat

Mooring              = rhadcp1_01_2020

WaterDepth           = 1080

MagDeviation         = -7.07

Start_Date           = 2020/10/15

Start_Time           = 19:30

End_Date             = 2022/07/20

End_Time             = 09:30

Latitude             = 57 36.89 N

Longitude            = 15 24.66 W

Columns              = z:instrument:serialnumber

1068	324	200044

1070	337	9378

## STAGE 1 – General Quality Control Checks (Nortek)

Wrapper: m_moorproc_toolbox\exec\gitrepo\stage1\adcp_s55\call_stage01_read_qc_S55_JC238.m

function: m_moorproc_toolbox\exec\gitrepo\stage1\adcp_s55\stage01_read_qc_S55.m

Should work with and without toolbox global variables set. Needed matlab toolboxes:

- m_moorproc_toolbox\exec\gitrepo\ needs to be added to matlab path for rodbload.m

- GSW v3.06 needs to be added to matlab path

Output: Data, figures and log-file

Data filename: moor_id_sn_stage1.nc,  e.g. for JC238: rhadcp_01_2020_200044_stage1.nc

Figures filenames: sn_stage1_f#_descriptor.png, e.g. for JC238:

- rhadcp_01_2020_200044_stage1_f1_pressure_tilt_heading_QC.png

- rhadcp_01_2020_200044_stage1_f2_beam_amplitude_correlation_QC.png

- rhadcp_01_2020_200044_stage1_f3_beam_amplitude_correlation_QC.png

- rhadcp_01_2020_200044_stage1_f4_velocity_and_speed_QC.png

- rhadcp_01_2020_200044_stage1_f5_horizontal_magnetometer_QC.png

- rhadcp_01_2020_200044_stage1_f6_sensor_diagnostics.png

Log file filename: moor_id_ADCP_stage1.nc,  e.g. for JC238: rhadcp_01_2020_ADCP_stage1.log

Data directory:  osnap\moorings\proc\moor_id\adcp_S55\

E.g. for JC238:  osnap\moorings\proc\rhadcp_01_2020\adcp_S55\

The function s55_load_nc_as_struct.m loads saved nc-file as matlab structure.

All checks should use plots showing original and flagged data. ONLY check and flag, processing is stage 2:

### Figure 1 – status: done

- Pressure: Must be reasonable; shallow readings may indicate deployment/retrieval and invalidate profiles. (deployment and retrieval automatically flagged)

- Instrument Orientation:

  - Tilt (Roll and Pitch) <10°: Ideal; data is within optimal design limits.

  - Tilt 10–30°: Compass and velocity errors increase—automatically flagged as QC_PROBABLY_BAD (3)

  - Tilt >30°:  Profiles likely erroneous (one transducer may be near-horizontal, making that velocity component undetectable)—automatically flagged as QC_BAD (4)

- Action:

  - Pressure: Verify readings are consistent with the deployment depth; data outside this range is automatically flagged.

  - Orientation: Tilt and Heading should remain stable. Large fluctuations suggest mooring instability or "knock-downs" by strong currents – unlikely for RHADCP.

  - Tilt 10–30°: Standard post-processing (e.g., bin-mapping) is not currently implemented. Development of these routines is only advisable if significant data gaps occur because of high-tilt periods.

### Figure 2 and Fig 3 – status: done

- Signal Level per Beam and Correlation:

  - Should start high, decrease exponentially with range.

  - Sudden increase at end of bin range = surface boundary; affected cells likely corrupted.

  - If no boundary, signal stabilizes; valid if above noise floor (correlation > 50%).

  - Cells near noise floor are invalid; measurements beyond this range are erroneous, i.e. all values with correlation <50% will be flagged as QC_BAD.

- Spikes in Signal Strength: Flag and remove affected cells

  - Mooring or Frame Interference

    - Appearance: Persistent "stripes" or constant high values at a specific depth in the amplitude plots (Figure 2).

    - Risk: Low, as the ADCP is top-mounted.

    - Action: Visually inspect Figure 2. Any constant horizontal stripe—outside of known surface bins—indicates a potential obstruction that requires manual flagging.

  - Transient Spikes (Fish, Bubbles, etc.)

    - Time-Burst Criterion: Calculates the Median and Median Absolute Deviation (MAD) for each 10-sample burst. Ensembles exceeding the burst median by 3 × MAD and a minimum 10 dB jump are identified as spikes.

    - Vertical Criterion: Monitors sudden amplitude changes along the depth profile. Bins deviating from their vertical neighbors by more than 3 × MAD and a minimum of 10 dB are flagged.

    - Logic: These "Loud" jumps (>10 dB, representing a 10-fold power increase) are flagged as QC_BAD, which ensures the corresponding velocity values are removed.

### Figure 4 – status: done

- Velocity Quality Control: Suspicious velocity values are flagged as QC_BAD (4) based on:

  - Spikes in horizontal velocity over depth (|ΔU|, |ΔV|) > 1.00 m/s

  - Temporal outliers > 3 standard deviations (per depth bin)

  - Vertical velocity outliers (|W|) > 0.10 m/s

### Figure 5 – status: done

- Compass Offset: Plot horizontal components as a circle.

  - Offset = “Hard Iron” rotating with instrument. See blue dots in figure 20.

  - Ellipse = “Soft Iron” nearby.

  - ONLY check that compass is resonable, hard iron compass correction is done in stage 2

Backround info - using the magnetometer data:

- Raw Measurement: The magnetometer outputs  My, Mx which contain the Earth's magnetic field distorted by the instrument frame's Hard-Iron error. The raw magnetic heading is calculated as: ψraw=atan2(-My, Mx)

- Calibration: By fitting a circle to the raw My, Mx data points, we identify the offsets yc,xc. The corrected heading relative to Magnetic North is found by centering the data: ψtrue=atan2-(My-yc), Mx-xc

- Correction Angle: The angular deviation caused by the Hard-Iron effect is:

ϵ=ψtrue-ψraw

- Velocity Transformation: To align the measured velocities Uraw, Vraw with the corrected heading, we apply a rotation matrix using the error ϵ:

Ucorr=Urawcosϵ-Vrawsinϵ

Vcorr=Urawsinϵ+Vrawcos⁡(ϵ)

The function s55_hard_iron_compass_correction.m calculates the center of the compass data (hard iron).In stage 1 we only plot it, but s55_hard_iron_compass_correction.m has the functionality to correct U and V accordingly, which can be used in stage 2.

### Figure 6 – status: completed

Plot all sorts of sensor diagnostics (temperature sensors, accelerometers, magnetometer). Action: Check then temperature and other sensors data are sensible, etc.

## STAGE 2 – Post-processing

Status: WIP

Wrapper: m_moorproc_toolbox\exec\gitrepo\stage2\adcp_s55\wrapper_stage02_proc_s55.m

Function: m_moorproc_toolbox\exec\gitrepo\stage2\adcp_s55\stage02_proc_s55.m

Should work with and without toolbox global variables set. Needed matlab toolboxes:

- m_moorproc_toolbox\exec\gitrepo\ needs to be added to matlab path for rodbload.m

- GSW v3.06 needs to be added to matlab path

- SCOTIA climatology needed: osnap\scotia\SCOTIA_monthly_clim_V8.nc

Output: Data, figures and log-file

Data filename: moor_id_sn_stage2.nc,  e.g. for JC238: rhadcp_01_2020_200044_stage2.nc

Figures filenames: moor_id_sn_stage2_f#_descriptor.png, e.g. for JC238:

- rhadcp_01_2020_200044_stage2_f1_hard_iron_correction.png

- rhadcp_01_2020_200044_stage2 _f2_UV_hard_iron_mag_dev_correction.png

- rhadcp_01_2020_200044_stage2_ f3_SoS_correction_range.png

- rhadcp_01_2020_200044_stage2 _f3_uv_first_bint.png

- rhadcp_01_2020_200044_stage2_ f4_sidelobe_interference.png

Log file filename: moor_id_ADCP_stage2.nc,  e.g. for JC238: rhadcp_01_2020_ADCP_stage2.log

Data directory:  osnap\moorings\proc\moor_id\adcp_S55\

E.g. for JC238:  osnap\moorings\proc\rhadcp_01_2020\adcp_S55\

The function s55_load_nc_as_struct.m loads saved nc-file from stage1 as matlab structure.

In this stage corrections are applied to data and the ensemble mean is calculated:

### Figure 1 – status: done

- Hard iron compass correction

This is the first correction. After plot comes up operator can select a time frame to improve circle fit. E.g. for JC238 ADCP rotated only at beginning.

### Figure  – status: done

- Correct horizontal velocity for Magnetic Deviation

After hard-iron compass and magnetic deviation correction is applied to U and V, the horizontal velocities are plotted. Please check rotations results in reasonable data

Ensemble mean DS_EM is now created – this means for Amplitudes and velocities all bad data are set NaN, so spikes/bad data do not impact ensemble mean. Any following data processing is applied to ensemble mean.

### Figure 3 – status: done

Uses SCOTIA data

- Calculate mid depth of each bin: 2 options available, figure is only displayed for option with MicroCAT

  - Option 1: Without MicroCAT data: depth from ADCP pressure and nominal range based on fixed SoS of instrument

    - TODO: Add option to the script asking for user provided mean SoS from previous data to get is at least a bit correct?

  - Option 2: Use MicroCAT data and SCOTIA – can be done during cruise using .use MicroCAT data - TODO: Add option to not use SCOTIA climatology in case it is not handy on board.

    - Uses MicroCAT pressure and correct for offset to ADCP sensor head

    - Uses MicroCAT T and S for speed of sound correction of bin ranges. Here three options are available:

      - MicroCAT+Scotia monthly climatology (time and depth varying)

      - MicroCAT only (time varying, constant with depth)

      - SCOTIA monthly climatology only ((time-interpolated profile, not recommended – only if MicroCAT data is bad maybe?– think about adding mean SoS on sensor head from previous data)

### Figure 4 – status: done

- Sde lobe interference is calculated and option to remove is given (recommended):

As RHADCP is mounted on string I applied the following including pitch.

Rmax=H cosθ0+θpitch-Cell Size

### Save data – status: outstanding

TODO: Save is as netcdf with the applied OceanSites convention (write external function for it and maybe adjust s55_load_nc_as_struct.m to also load STAGE2 data accordingly

## STAGE 3 – Delayed time post-processing

Run stage 2 again with calibrated MC data

## Appendix

Other notes just to have some equations:

Calculate depth from surface signal in beam amplitude:

- Quadratic fit over surface bins to determine bin fraction with highest echo amplitude defining the surface

- Calculate nominal bin distances to surface (bin-1)*cell_size - (i.e. distance of first bin to first bin is 0, second bin is 1*cell_size, etc) - interpolate nominal bin distances on surface bin fraction - This give you the distance of the first bin to surface (uncorrected) hi (for each beam)

- Correct for speed of sound mean climatology profile at ADCP location (ceff: harmonic mean over all depth - use pressure sensor here, cinstr: set to 1500m/s):

hic=ceff/cinstr*hi

- Then they add the distance from first bin to transducer

- From that the correct for the beam width, accounting for transmission loss:

  - EL=SL-2αR-20log10(R)

  - What is absorbtion coeff α for 55kHz?

  - What is the beam slant angle for 55kHz? - 20°?

  - Surface roughness parameter ( aether2 ) is computed from wind speed at the mooring:

aether2=0.003+5.12×10-3×wspadcp

- For each time step:

Ss=-10*log8 π aether2-2.17aether2 tan(90-ϴ)360 2 π2

- The reflected echo level is: EL = SL + TL + S_s

- The beam angle with maximum echo level defines the effective angle:

θeff = mean(θmaxEL)

- For each beam ( i ) and time step:

hib=cosθeff360 2πcosθ0360 2πhi	, θ0 = beam slant angle - 20° for S55?

- Correct for pitch and roll - need new code for 3 beam solution and measured pitch and roll:

θeff =arccoscosθ0-sinϕp*sinθ0*cosψi+sinϕr*sinθ0*sinψi

hcorr= hib*cosθeffcosθ0

Where:

θ0= nominal beam angle

ϕp= pitch rotation

ϕr= roll rotation

ϕi= azimuth angle of the beam (e.g., 0°, 120°, 240° for beams 1–3)

- Estimate depth offset: compare mean depth derived from pressure sensor with mean depth derive from surface amplitude